Direct Monitoring of Individual Ammonia or Volatile Fatty Acids Using a Temperature-Controlled Galvanic Sensor Chip

Farahdiana Wan Yunus, Moataz Mekawy, Edhuan Ismail, Jin Kawakita,* and Izumi Ichinose*

Research Center for Macromolecules and Biomaterials, National Institute for Materials Science, 1-1 Namiki, Tsukuba, Japan

*Corresponding authors: KAWAKITA.Jin@nims.go.jp, ICHINOSE.Izumi@nims.go.jp

Abstract

The detection of ammonia (NH3) and volatile fatty acids (VFAs) during anaerobic processes plays an important role in the management of livestock and their waste, as well as in biogas production. A newly developed moisture sensor chip (MSC) was used to monitor these gasses volatized from aqueous solutions. The concentrations of NH3, acetic acid (AA), propionic acid (PA), and butyric acid (BA) were in the range 60–150 mM. To substantially improve the detection sensitivity, MSC was cooled from 24 to 19 °C. Because the temperature was decreased, volatized water and solute condensed on the surface of the MSC and the formed microdroplets and the detection current increased to several thousand picoamperes. The detection current also increased with increasing concentration of VFAs and was highly sensitive to the pH conditions and the partial pressure of VFAs in the air phase. The applicable concentration range was from 0 to 150 mM for NH3 and from 60 to 150 mM for VFAs. The highest sensor sensitivity to NH3 was 12.3 mM-1 at 22 °C. The highest sensor sensitivity for AA among VFAs was 4.0 mM-1 at 19 °C. The amount of water droplets formed on the surface of MSC increased with increasing pH of the aqueous solution in the order NH3>water>VFAs. Droplet formation preferentially occurred in the order AA>PA>BA, indicating that a shorter alkyl chain of VFA was preferred. The proposed temperature-controlled MSC is applicable for the detection of NH3 under basic conditions and for the detection of VFAs under acidic conditions.

Keywords: Ammonia (NH3), volatile fatty acids (VFAs), acetic acid (AA), propionic acid (PA), butyric acid (BA), moisture sensor chip (MSC)

- Introduction

Ammonia (NH3) is a critical material in many industrial processes, and it is being considered as a future fuel in power plants [1] and a precursor of hydrogen fuel [2]. NH3 is also produced widely via anaerobic processes in landfills [3], sewage treatment plants [4], and agricultural activities [5–7]. However, NH3 concentrations that exceed the safe range can cause numerous human health and environmental problems. Access to safe water is one of the 17 sustainable development goals (SDGs) outlined by the United Nations. Water polluted with NH3 harms aquatic ecosystems and is not potable. Long-term exposure to NH3 gas causes health issues such as eye irritation, and damage to vital organs [8]. Available methods for monitoring NH3 levels in water require a liquid sampler to form gaseous samples suitable for gas chromatography [9,10]. Chemical indicators and electrochemical probes for NH3 are available. However, the sensitivity of chemical sensors decreases with the consumption of chemicals. One difficulty in NH3 detection is that various nitrogen compounds coexist in the aquatic environment and their concentrations vary depending on other chemicals. Wang et al. pointed out that water pollution is a complex issue and needs a simple real-time monitoring for nitrogen-based compounds [11]. Japan's national Science and Technology Research Partnership for Sustainable Development (SATREPS) is collaborating with developing countries to control the levels of nitrogen-based water pollutants [12].

The same anaerobic processes that produce NH3 also produce VFAs [13,14]. In particular, large amounts of VFAs are produced in the digestion system of ruminants, where microorganisms decompose the organic materials into various carboxyl compounds and alcohols [15]. Like NH3, VFAs are important intermediates or products in industries such as biogas production [16], food additives [17], and chemical manufacturing [18]. In ruminants, which represent a vital food source for people, VFAs are metabolized to provide energy [19]. Giardi et al. [20] and Dinsdale et al. [21] have noted that VFAs serve as an important indicator for ruminants’ health and for dairy and meat production. Among ruminants, cows produce the largest amount of VFAs [22]. When the VFA concentration becomes too high, the resultant low-pH condition leads to subacute ruminal acidosis, abscess of the liver, and inflammation of tissue above the hoof [23–25]. For these reasons, the development of an NH3 and VFA monitoring system is critical for the livestock and agricultural sectors [26]. Simple monitoring of pH cannot estimate the activity of a microbiome. The productions of NH3 under high pH and that of VFAs under low pH need to be monitored.

Conventionally, NH3 has been measured using microdiffusion adsorption, ion-selective electrodes, and titration [27]. The concentration of NH3 could be measured using the colorimetric method or microdiffusion. Meanwhile, the total or individual VFA concentration has been measured using chromatography techniques [28–31]. Korent et al. achieved the electrochemical detection of NH3 in a biological fluid using an Au electrode coated with polyaniline that undergoes a change in oxidation state depending on the concentration of NH3 [32]. Minami et al. demonstrated the strong potential of a nanomechanical sensor array that can distinguish silage gas samples containing different concentrations of VFAs at parts-per-million concentrations [33]. Lamb et al. have developed an organic dye-based sensor that is promising for detecting low concentrations of VFAs via a color change [34]. Recently, Murata et el. developed a highly selective organic gas sensor using DNA as a biomaterial [35]. To the best of our knowledge, there are no studies on the simultaneous detection of NH3 and VFAs using a single sensor chip.

Herein, we studied the sensing ability of humid gases volatized from aqueous solutions of NH3 and VFAs by measuring the galvanic current from our MSC [36]. MSC could detect adsorbed water molecules with high sensitivity and monitor the condensation behavior [37]. NH3 and VFA vapors co-adsorb with water molecules onto the MSC surface. Because the MSC is not directly exposed to aqueous samples, this detection method can prevent the corrosion of electrodes and maintain high sensitivity [38]. The aim of the present work is to monitor the adsorption and condensation behaviors of aqueous solutions containing different concentrations of NH3 and VFAs. The condensation behaviors were studied alongside microscopic observations. Careful analysis of the behaviors revealed that, in addition to the concentrations of volatile solutes (i.e., NH3 and VFAs), the change in surface wettability also determines the response currents. This newly developed microsensor can be used for monitoring microbial activities under a wide range of anaerobic conditions such as cow rumen.

- Materials and method

2.1.	Materials

PA (> 99.5% purity) and BA (>99.0% purity) were purchased from TCI, and AA (99.9% purity) and 28 wt% NH3 aqueous solution were obtained from FUJIFILM Wako Chemicals. Aqueous solutions with a concentration range from 60 to 150 mM were prepared. A wide concentration range was chosen to include the maximum values reported for silages, anaerobic digester, and cow rumen (Table S1). Concentrated NH3 or VFA gases are not available.

2.2.	Experimental section

2.2.1.	Sensor configuration

The fabrication of the MSC has been explained elsewhere [39].  Schematics of the sensor configurations are shown in Figure 1a and 1b. Sensor response was measured as galvanic current between two electrodes of Al and Au by using a tailored module with a current-input A/D converter (DDC118, Texas Instruments). The area of the chip is 0.25 cm2. Note that Al electrode is not very stable for strong acidic or strong basic conditions. We have not experienced any corrosion while detecting moisture. Improvement of Al electrode is now under our consideration. For conditions described later, a single sensor was used at least 3 times to detect the NH3 and VFAs. We also repeated the same experiment with different sensors and confirmed that the detection tendency was the same with very minor changes. Sensor’s long-term stability has been confirmed at least in our experimental conditions.

Fig. 1. (a) Schematic of the MSC configuration. (b) Cross-sectional configuration. (c) Experimental setup. (d) Photos of the actual setup.

    - Measurement and observation procedures

A schematic and photos of our experimental setup is shown in Figure 1c and 1d. MSC was placed inside a small chamber where its temperature could be controlled between 24–19 °C by attaching a Peltier device to its bottom. The small chamber (40 × 38 × 13 mm) and buffer chamber were placed inside a large container with fixed temperature of 25 °C. An NRP 3000-EYELA pump was used to circulate the NH3, or VFAs vapors into the buffer and the MSC chamber.

The vapors slowly condense into liquid on the MSC surface when its temperature is adjusted to be lower than the chamber temperature. The condensation of vapors from NH3 or VFAs on MSC surface leads to a large increase of the detected current. We confirmed the correlation of MSC response and measured relative humidity (RH) by a hygrometer. At 25 °C, RH values were <40% for water, 150 mM VFAs, or NH3. RH became 100% when temperature decreased to 20 °C, and then the current signal was recorded. The response time (TRes) is the time required by MSC to reach 90 % of saturated current response. The recovery time (TRec) is the time taken by MSC to reach back to 10%. We stopped the supply of  the air containing VFAs or NH3 by removing the vial and flowed fresh air to dry the sensor while measuring the sensor response as a detected current. The sensor response (IRes) was calculated by using the following equation:

IRes=It, c - I0I0 = ∆II0

, where It,c is the detected current at a certain time and concentration, I0 is the initial detected current in dry air (~5 pA) [40]. The sensor response at each temperature (24, 22, and 19 °C) was plotted against the concentration of each solution. The sensor sensitivity for concentration was calculated from the slope of the concentration-dependence relationship with the response. Finally, the lower level of detection (LOD) in mM for each solution at certain temperature was calculated by the following equation:

LOD=3.3×SyS

, where S [mM-1] is the sensitivity and Sy is the standard deviation.

A Keyence VHX digital microscope was used for the microscopic observation of MSC surface. The optical images were analyzed using the ImageJ software to calculate the droplets' coverage of the MSC surface. Contact angle measurements were carried out using a Kyowa DropMaster to check the MSC surface wettability. MSC was placed on the stage of the apparatus, and then a 1.0 μL droplet of the tested solutions was placed on the surface. The values were obtained from the average of five measurements. The pH was measured using a pH meter from TOKO.

3.	Results and discussion

  - Condensation behavior of NH3 and VFA solutions

Fig. 2a shows the response of the MSC exposed to vapor generated from 150 mM NH3, 150 mM PA, 150 mM AA, and 150 mM BA solutions and water as a reference. The bar chart was replotted from Fig. 2a by averaging last 5 minutes at each temperature (Fig. S1a). The sensor response for 150 mM NH3 increased to 275 at 24 °C in 16 min. When the temperature was lowered to 19 °C, the response increased to 3500, which was the maximum limit of our sensor. Similar finding was obtained for VFAs and water. From Fig. 2a, we concluded that the 150 mM NH3 solution gave far higher responses than VFAs and water. Detailed explanations were given in the supporting information.

The 150 mM AA showed higher response than water at all temperatures except at 19 °C. There are two reasons, 1) AA becomes less conductive than water. At temperatures lower than 23 °C, AA generates ionic species in the thin water layer. 2) The increase in hydrophobicity of the MSC surface. This effect prevents the accumulation of water molecules at the surface. The 150 mM BA solution showed the lowest response among the VFA solutions. The strong hydrophobic effect of BA must prevent the accumulation of water even at low temperatures [41]. The detected response signals from MSC rely on the concentration, temperature, and surface wettability. As explained in the next section, the sensitivity towards each gas differs by temperature. Detection current increased at highly humid conditions is the strength of MSC (Table S2). The strength of MSC is the fact that the detected current could increase at highly humid conditions (Table S2).

The microscopic images reveal that no water droplets were present at 22 °C, indicating that NH3 and VFA molecules are in a thin layer of adsorbed water in the temperature range 24–22 °C (Figure 2c–g (i)). At 21 °C, in the NH3 and VFAs cases, water droplets started to be observed on the MSC surface (Fig. 2c–d (ii) and 2f–g (ii)). By contrast, for water, condensation started at 20 °C (Fig. 2e (iii)). Droplet coverage for NH3 solutions increased from 1% (21 °C) to 49% (19 °C). Droplet coverage for PA solutions increased from 5% to 30% at these temperatures. Water gave 3% and 35% coverages at 21 °C and 19 °C, respectively. Among the three investigated VFAs, AA showed the largest droplet coverage of 34% at 19 °C. This coverage decreased with increasing molecular weight of the VFAs (30% for PA and 5% for BA). The details of the values are also described in Table S3, and the droplet coverage traced values are shown in Figure S2.

Fig. 2. (a) The response of the MSC to vapors generated from water, NH3 and VFA solutions, plotted as a function of time. The temperature was varied with time from 24 to 19 °C. (b–f) Microscopic images of the surface of the MSC at temperatures from 22 to 19 °C for the vapors generated from 150 mM NH3, 150 mM PA, 150 mM AA, and 150 mM BA solutions and from water.

  - Concentration-dependent response

The concentration dependency of the sensor response was studied at three temperatures that correspond to the initial (24 °C) and middle (22 °C) stages of the formation of a water layer with a uniform thickness and the final stage of droplet accumulation (19 °C). The results are analyzed as shown in Fig. 3a and summarized in Fig. 3b–e for each solution; the raw data are shown in Figure S1c–e. As the concentrations of the NH3 and VFAs were increased from 60 mM to 90 mM, 120 mM, and 150 mM, the initial-stage (24 °C) response of the NH3 solutions increased substantially from 279 (60 mM) to 631 (150 mM). The VFAs also showed a concentration dependency at this stage. The response of the PA solutions increased from 11 (60 mM) to 25 (150 mM). AA exhibited smaller response of 4 (60 mM) and 5 (150 mM), and the BA response slightly increased from 3 (60 mM) to 4 (150 mM). Compared with the response of 7.6 for pure water, that of NH3 is very high; however, those of AA and BA are small.

During the middle stage (22 °C), the response to water increased to 63. In sharp contrast, NH3 solutions gave large response of 849 (60 mM) to 2023 (150 mM), which was 13–32 times the response of water. The VFAs showed small increases of the response with increasing concentration. However, the BA solutions showed lower values than water. Specifically, the responses were 71 (60 mM) to 91(150 mM) for PA, 11.6 (60 mM) to 139 (150 mM) for AA, and 13.6 (60 mM) to 16 (150 mM) for BA. At this temperature, only NH3 and AA showed a clear concentration dependency, as indicated by Fig. 3b, c. Our results at 22 and 24 °C confirmed that the linearity in response from 0 to 150 mM could be achieved only with NH3. Similarly, a linear response for AA was found at 22 °C from 60 to 150 mM.

At the final stage (19 °C), the NH3 solutions reached the maximum response of our sensor system, giving a value of 3500 for all concentrations. Unlike NH3, PA showed a decrease in response as the concentration was increased from 60 mM (749) to 150 mM (381). This tendency was also observed for BA, for which the response decreased from 589 (60 mM) to 249 (150 mM). By contrast, AA showed an increase of the responses, where the values ranged from 647 (60 mM) to 1005 (150 mM). The difference in VFAs responses indicates that the degree of wettability of the MSC surface toward these aqueous solutions differs. All of these experimental data are summarized in Table S4.

Figure 3f shows the sensitivity of the sensor (mM-1) for each solution as a function of temperature. At 22 °C, the sensor showed the highest sensitivity to NH3 (12.3/mM) followed by 3.8/mM at 24 °C. No significant sensitivity was observed for NH3 at 19 °C although it gave the highest but fixed response. In case of VFAs, the sensor showed its higher sensitivity to AA (however it is lower than that of NH3) at 22 and 19 °C. The lower level of detection (LOD) for each solution was calculated at each temperature as shown in Fig. 3g. At 24 °C, the LOD order was NH3>BA>AA>PA. At 22 °C, the order was AA>NH3>PA>BA. At 19 °C, only AA showed a positive LOD among VFAs. The response and recovery times for the sensor were evaluated at 19 °C where the highest droplet coverage occurred (Fig. 3h). NH3 showed the fastest response at ~12.3 min. This followed by BA, PA, and AA at 21, 24, and 27 min, respectively. BA has the fastest recovery at 1.5 min followed by AA, PA, and NH3 at 6.5, 7.5, and 8 min, respectively. Cross-sensitivity measurements and confirmation of the long-term reproducibility are in progress in our laboratory.

We comprehensively compared our sensor with previously reported sensors and listed the parameters in Table S5. Previous studies concluded that the applicable concentration ranges were from 0 to 250 mM for VFAs and from 0 to 7700 mM for NH3. LOD values were about 270 M for VFAs and 3×10-6 M for NH3, respectively. These ranges and values are based on the collective study of many sensors that could only detect either VFAs or NH3, separately. The concentration ranges of our sensor are from a single sensor that could detect both VFAs and NH3. Our study achieved the linear detection range from 60 to 150 mM for VFAs and from 0 to 150 mM for NH3. LOD value was 80 M for VFAs (AA) and 900 M for NH3.

Our MSC is a galvanic sensor which functions based on the condensation of the target gases from very humid conditions usually found in sewage, digesters, and cow rumen. It is highly sensitive, and concentration changes are detected in the order of picoamperes. In terms of the detection of NH3, our sensor could detect it at high concentrations but at a longer response time. The response time is comparable to a bioelectric sensor for VFA detection (see references in Table S5).

Fig. 3i–m shows optical microscopic images of droplets on the MSC surface at the final stage (19 °C). The large coverage of the droplets observed (except for the BA solutions) indicates that condensation has been fully proceeded. The droplet images of the NH3 solution and water are similar, showing droplets widely spread on the MSC surface. Compared with the droplet size for water (Figure 3i), the droplet sizes for the NH3 solutions are slightly larger (Fig. 3j). Compared with NH3 and water, the VFAs showed smaller droplets. Among the VFAs, PA exhibited the largest droplet size of approximately 136–149 m in our investigated concentration range (Fig. 3k). The AA droplets in Fig. 3l show a clear increase in droplet size as the concentration increased. The droplets increased from 70 m (60 mM) to 134 m (150 mM), indicating an increase in hydrophilicity of the MSC surface. In contrast to the droplet size of AA, that of BA decreased substantially with increasing concentration and no droplets larger than 8 m were obtained at a concentration of 120 mM or more (Fig. 3m). This result indicates that the surface becomes hydrophobic with increasing BA concentration [41,42].

Fig. 3. A typical current response and determination of response and recovery time (a). Concentration-dependent response for NH3 and VFA solutions (b-d) at 24 °C (initial stage, black), at 22 °C (middle stage, red), and 19 °C (final stage, green). Sensor sensitivity (e), limit of detection (LOD) (f) for all solutions at 24, 22, and 19 °C. Response and recovery times (g) for all solutions during and after full wettability at 19 °C. Optical microscopic images for the droplets of  water (h) and  NH3 (i), PA (j), AA (k), and  BA (l) solutions formed on the MSC surface at 19 °C. Droplet size was monitored at concentrations of 60–150 mM. Photos were obtained at 90 min experimental time.

The large expansion of the NH3 solution droplets is due to the deprotonation of silanol units on the silicon dioxide (SiO2) surface between the Au and Al electrodes of the MSC. The high vapor pressure of NH3 enables it to easily accumulate on the MSC surface with water vapor; however, the transport rate of water vapor might slightly decrease in the presence of 150 mM NH3 solution (Fig. 3j). By contrast, the droplet size increased with increasing AA concentration (Fig. 3l), which can be explained by the increase in hydrophilicity of the MSC surface upon protonation of the SiO2 surface by AA. The coordination of AA to the SiO2 surface leads to the observed increase in hydrophobicity; however, this effect appears to be weak. The effect of protonation is consistent with the observation that the sensor response increased with increasing concentration of AA (Fig. 3c). The droplets formed from PA solutions were approximately the same size irrespective of the PA concentration (Fig. 3k). We speculate that the protonation effect (i.e., an increase in hydrophilicity) and the coordination effect (i.e., an increase in hydrophobicity) are almost compensated [43]. Droplets generated from BA solutions decreased in size with increasing BA concentration (Fig. 3m). The MSC surface becomes more hydrophobic with increasing BA concentration because of its longer alkyl chain [44,45]. Our results demonstrate that the order in which aqueous solutions accumulate on the MSC surface is NH3 > water > AA > PA > BA. These findings were confirmed by contact angle measurements and can be explained by the formation of a cationic or anionic MSC surface, as described later.

  - Droplet coverage on MSC surface at 19 °C

Detailed observations of droplet formation at different exposure times were conducted to explain the detected current behaviors of the MSC toward water and NH3 and VFA solutions. Figure 4(a–e) shows the microscopic images after the temperature was lowered to 19 °C (76.5–84 min). Fig. 4f shows the coverage of droplets on the MSC surface, corresponding to Fig. 4a–e. Coverage is described as percentages in Fig. 4f and is summarized in Table S6; the ImageJ raw data are shown in Figure S3. The results were repeated (Figure S1b).

At the beginning of droplet formation (stage i), the coverages for NH3, PA, water, and AA cases were no more than 3.6%, 1.6%, 2.4%, and 1.7%, respectively. Meanwhile, droplet coverage remained constant at 0.2% for BA. The droplet coverage for NH3 was 1.5 times greater than that for water, which was 1.5 times greater than that for PA. In the subsequent 3 min (stage ii), the droplet coverage for NH3 increased to 10% and those for PA, water, AA, and BA were 2.6%, 4.4%, 3.9%, and 0.4%, respectively. The coverage increase for AA is relatively greater than that for PA. During the last 3 min (stage iii), the droplet coverages were 22%, 4%, 18%, 17%, and 0.9% for NH3, PA, water, AA, and BA, respectively. In general, NH3, water, and AA gave higher coverages, whereas PA and BA exhibited lower coverages. These coverage values are consistent with the sensor response in Fig. 2a.

Fig. 4.  (a–e) Optical microscopic images of the droplet coverage in the time range between 76.5 and 84 min for 150 mM NH3, 150 mM PA, water, 150 mM AA, and 150 mM BA. (f) Droplet coverages on the MSC surface at the final stage (19 ℃). (g) Changes of the droplets coverages as a function of the detected current. Currents were monitored at stage (i) 76.5–78 min, stage (ii) 79–81 min, and stage (iii) 82–84 min.

Detected currents at all of the coverages shown in Fig. 4a–e were measured; the results are plotted in Fig. 4g as a function of droplet coverage. NH3 showed the highest current of 17,500 pA for the range of droplet coverage, and this value was constant. AA also showed constant currents of ~4000 pA, presumably because of the high ionic conductivity of the thin water layer. By contrast, water showed an increase in current when the droplet coverage was small and then reached a steady state at 5000 pA. PA and BA showed currents that linearly increased to 2000 pA and 1000 pA, respectively, with increasing droplet coverage. These results indicate that the droplets contribute to the conductivity. Because of their hydrophobic alkyl chain, droplet formation on the MSC surface was more feasible than on the AA surface. As shown in Ref. [46], carboxylic acid undergoes a condensation reaction to form ester bonds with the hydroxyl group on the SiO2 surface:

R-COOH+HO-Si→R-COO-Si+H2O.

As the alkyl chain (R–) increases in length, the hydrophobic effect becomes stronger. BA, which has the longest alkyl chain among the investigated VFAs, can therefore not form large droplets even at 19 °C.

  - MSC surface wettability

MSC surface wettability is one of the most important factors for the detection of NH3 and VFAs in water. Therefore, we carefully measured the contact angles of the MSC surface, changing the temperature of these solutions. As evident in Fig. 5a, the contact angles increased by approximately 6–8° for all of the solutions when the MSC temperature was lowered from 24 to 19 °C. The measured values were 61°–69° (NH3), 67°–75° (PA), 65°–73° (water and AA), and 70°–76° (BA). Details of the contact angle measurements are presented in Figure S4. The order of the contact angles is NH3<water≈AA<PA<BA. The surface tension of an aqueous solution is known to increase with decreasing temperature [47]. Our observations reveal a similar tendency. Interestingly, BA, which has a longer alkyl chain than the other investigated VFAs, showed the largest contact angle. In general, hydrophobic molecules lower the surface tension of water, like surfactant molecules, and lead to a decrease in the contact angle. The fact that BA exhibited the highest contact angle strongly suggests that the hydrophobicity of the MSC surface increased substantially and that this effect was more significant than the decrease in the surface tension of the BA solution. PA showed a similar effect (i.e., it coordinates to the SiO2 surface and makes the surface hydrophobic). NH3 did not decrease the surface tension of the aqueous solutions. However, we observed contact angles lower than that of water. This result indicates that the MSC surface becomes more hydrophilic because of deprotonation of the silanol groups. AA has a contact angle similar to that of water, indicating that the protonation effect and the decrease in the surface tension are compensated. As evident in Figure 5b, the detected currents are correlated to the contact angle values. The NH3 solutions had the smallest contact angles among the investigated solutions and exhibited the highest currents of 1000–16,000 pA. BA exhibited the highest contact angle, which led to it exhibiting the lowest detected currents of 30–1100 pA. PA showed a similar tendency as BA, exhibiting relatively high contact angles and lower currents of 100–1800 pA. The contact angles and detected currents (60–4000 pA) of water and AA show similar tendencies.

Fig. 5. (a) Changes in contact angles as a function of temperature. The concentrations of NH3 and VFAs are 150 mM. (b) Changes of detected currents as a function of contact angle. (c) Changes of the contact angle at different pH levels and temperatures. PZC: point of zero charge (for SiO2). (i) acidic range, (ii) PZC pH, and (iii) basic range.

Figure 5c shows a plot of the contact angle as a function of pH. The values at pH 10.9 correspond to a 150 mM NH3 solution. The contact angles near pH 2.7 correspond to 150 mM VFA solutions. The color changes indicate the temperature in the range from 24 °C (red) to 19 °C (blue). As the temperature decreased, the contact angles of solutions increased because of the reduced surface tension. These changes can be discussed on the basis of the point of zero charge (PZC) of the SiO2 surface. SiO2 has a PZC value of pH 2.8 [48, 49]. At this pH, SiO2 has no overall surface charge, which is denoted by (ii) in Figure 5c. VFA solutions with a pH near 2.7 may slightly protonate the SiO2 surface, forming Si–OH2+ and thereby imparting a slight positive charge. This range is denoted by (i) in Figure 5c. Conversely, water and NH3 (pH 10.9) solutions, which have pH levels far greater than the PZC of the SiO2 surface, will enhance deprotonation, forming Si–O− on the surface and thereby increasing the amount of negative charge, as shown by range (iii) in Fig. 5c. As reported by Virga, the solution wettability increases with increasing amounts of ions on the metal oxide surfaces [50]. Notably, the BA, PA, and AA solutions exhibit different contact angles despite their pH levels being similar to the PZC value of SiO2. These different contact angles indicate that the coordination effect is far more significant than the increase in positive charge.

- Conclusions

A newly developed MSC was applied to the detection of NH3 and VFA vapors for the first time. The results revealed that, among the investigated analytes, NH3 was detectable with the highest sensitivity. The large response for NH3 is due to the deprotonation of the silanol groups on the MSC surface. This reaction increases the ion concentration and conductivity of the thin water layer, resulting in a large increase of the MSC surface wettability and in a detected current of a few thousand picoamperes. VFAs were also detectable when an appropriate temperature was selected for the MSC. The length of VFAs' alkyl chain changes the hydrophobicity of the MSC surface, leading to MSC response that decreases in the order AA>PA>BA. BA exhibits the smallest response because the MSC cannot form the water droplet even at 19 °C. That is, the concentration of BA is determined by the decrease in the current, especially at low temperatures. The MSC is a promising sensor with a chip size of 0.25 cm2 and a high sensitivity of thousands of picoamperes. The detection rate of our sensor at 90% response is ~12.5 min for NH3 and ~25 min for VFAs. We are now developing a method to bring better MSC selectivity towards simultaneous detection of NH3 and VFAs. Some examples for future practical applicability may be extended to detect the gases from cow rumen, sewage, bio-gas production processes, and fermentation processes. The presence of both NH3 and VFAs in these environments has been confirmed and our sensor is suitable for detecting these analytes.

CRediT authorship contribution statement

F. W. Yunus: Methodology, Validation, Investigation and Formal analysis, Data curation, Writing – original draft, M. Mekawy: Validation, Investigation and Formal analysis, Writing - review & editing. E. Ismail: Investigation and Formal analysis, Writing - review & editing. J. Kawakita: Conceptualization, Methodology. I. Ichinose: Funding acquisition, Writing - review & editing, Supervision.

Acknowledgement

This work was funded by the MOONSHOT R&D program JPJ009237 and partially supported by JSPS KAKENHI Grant Number 22H02149.

References

[1] H. Tian, K. Liu, J. Hao, Y. Wang, J. Gao, P. Qiu, C. Zhu, Nitrogen oxides emissions from thermal power plants in China: Current status and future predictions, Environ. Sci. Technol. 47 (2013) 11350–11357. https://dx.doi.org/10.1021/es402202d

[2] S. Giddey, S.P.S. Badwal, C. Munnings, M. Dolan, Ammonia as a renewable energy transportation media, ACS Sustainable Chem. Eng. 5 (2017) 10231–10239. https://doi.org/10.1021/acssuschemeng.7b02219

[3] Y.Y. Long, Y. Liao, J.Y. Miao, D.S. Shen, Effects of ammonia on methane oxidation in landfill cover materials, Environ. Sci. Pollut. Res. 2 (2014) 911–920. https://doi.org.10.1007/s11356-013-1963-8

[4] X. Shi, J. Lin, J. Zuo, P. Ji, X. Li, X. Guo, Effects of free ammonia on volatile fatty acid accumulation and process performance in anaerobic digestion of two typical bio-wastes, J. Environ. Sci. (China) 55 (2017) 49–57. http://dx.doi.org/10.1016/j.jes.2016.07.006

[5] R.B. Bist, S. Subedi, L. Chai, X. Yang, Ammonia emissions, impacts, and mitigation strategies for poultry production: A critical review, J. Environ. Manage. 328 (2023) 116919 https://doi.org/10.1016/j.jenvman.2022.116919

[6] L. Wang, D. Huang, Soil ammonia-oxidizing archaea in a paddy filled with different irrigation and fertilization managements, Sci. Rep. 11 (2021) 14563. https://doi.org/10.1038/s41598-021-93898-y

[7] A. Parodi, W.J.J. Gerrits, J.J.A.V. Loon, I.J.M.D. Boer, A.J.A. Aarnink, H.H.E.V. Zanten, Black soldier fly reared on pig manure: Bioconversion efficiencies, nutrients in the residual material, greenhouse gas and ammonia emissions, Waste Management 126 (2021) 674–683. https://doi.org/10.1016/j.wasman.2021.04.001

[8] X. Wang, L. Gong, D. Zhang, X. Fan, Y. Jin, L. Guo, Room temperature ammonia gas sensor based on polyaniline/copper ferrite binary nanocomposites, Sens. Actuators B: Chem. 322 (2020) 128615. https://doi.org/10.1016/j.snb.2020.128615

[9] K. Hayashi, M. Mano, K. Ono, T. Takimoto, A. Miyata, Four-year monitoring of atmospheric ammonia using passive samplers at a single-crop rice paddy field in central Japan, J. Agric. Meteorol. 69 (2013) 229–241. https://doi.org/10.2480/agrmet.69.4.1

[10] A.A. Puchalski, M.E. Sather, J.T. Walker, C.M.B. Lehmann, D.A. Gay, J. Mathew, W.P. Robarge, Passive ammonia monitoring in the United States: Comparing three different sampling devices, J. Environ. Monit. 13 (2011) 3156–3167. https://doi.org/ 10.1039/c1em10553a

[11] M. Wang, A.B.G. Janssen, J. Bazin, M. Strokal, L. Ma, C. Kroeze, Accounting for interactions between sustainable development goals is essential for water pollution control in China, Nat. Commun. 13 (2020) 730. https://doi.org/10.1038/s41467-022-28351-3

[12] A. Popova, R. Rattanakom, Z.Q. Yu, Z. Li, K. Nakagawa, T. Fujioka, Evaluating the potential of nanofiltration membranes for removing ammonium, nitrate, and nitrite in drinking water sources, Water Res. 244 (2023) 120404 https://doi.org/10.1016/j.watres.2023.120484

[13] M.L.U. Rehman, A. Iqbal, C.C. Chang, W. Li, M. Ju, Anaerobic digestion, Water Environ. Res. 91 (2019) 1253–1271. https://doi.org/10.1002/wer.1219

[14] F. Bonk, D. Popp, S. Weinrich, H. Strauber, S. Kleinsteuber, H. Harms, F. Centler, Ammonia inhibition of anaerobic volatile fatty acid degrading microbial communities, Front. Microbiol. 9 (2018) 2921. https://doi.org/10.3389/fmicb.2018.02921

[15] A. Rotbart, P.J. Moate, C.K. Yao, J.Z. Ou, K.K. Zadeh, A novel mathematical model for the dynamic assessment of gas composition and production in closed or vented fermentation systems, Sens. Actuators B: Chem. 254 (2018) 354–362. http://dx.doi.org/10.1016/j.snb.2017.07.066

[16] H.W. Schroer, C.L. Just, Feature engineering and supervised machine learning to forecast biogas production during municipal anaerobic co-digestion, ACS EST Engg. 4 (2024) 660–672 https://doi.org/10.1021/acsestengg.3c00435

[17] S. Agnihotri, D.M. Yin, A. Mahboubi, T. Sapmaz, S. Varjani, W. Qiao, D.Y. Koseoglu-Imer, M.J. Taherzadeh, A glimpse of the world of volatile fatty acids production and application: A review, Bioengineering 13 (2022) 1249–1275. https://doi.org/10.1080/21655979.2021.1996044

[18] M. Atasoy, I. Owusu-Agyeman, E. Plaza, Z. Cetecioglu, Bio-based volatile fatty acid production and recovery from wastestreams: Current status and future challenges, Bioresour. Technol. 268 (2018) 773–786. https://doi.org/10.1016/j.biortech.2018.07.042

[19] P. Cholewinska, K. Czyz, P. Nowakowski, A. Wyrostek, The microbiome of the digestive system of ruminants – A review, Anim. Health Res. Rev. 21 (2020) 3–14. https://doi.org/10.1017/S1466252319000069

[20] S. Silletti, G. Rodio, G. Pezzotti, M. Turemis, R. Dragone, C. Frazzoli, M.T. Giardi, An optical biosensor based on multiarray of enzymes for monitoring a large Set of chemical classes in milk, Sens. Actuators B: Chem. 215 (2015) 607–617. http://dx.doi.org/10.1016/j.snb.2015.03.092

[21] R.F. Feito, R.M. Dinsdale, A.J. Guwy, G.C. Premier, Applicability of a PEDOT coated electrode for amperometric quantification of short chain carboxylic acids, Sens. Actuators B: Chem. 225 (2018) 712–719. https://doi.org/10.1016/j.snb.2017.08.033

[22] J. Dijkstra, J. M. Forbes, J. France, Quantitative Aspects of Ruminant Digestion and Metabolism, second ed., CABI Publishing, United Kingdom, 2005, pp. 263–288.

[23] J.B. Russell, Mutualism and commensalism / rumen, in: Encyclopedia of Microbiology, Vol. 1, Elsevier, San Diego, 2009, pp. 163–174.

[24] E. Khafipour, S. Li, J.C. Plaizier, D.O. Krause, Rumen microbiome composition determined using two nutritional models of subacute ruminal acidosis, Appl. Environ. Microbiol. 75 (2009) 7115–7124. https://doi.org/10.1128/AEM.00739-09

[25] J. Zosel, H. Kaden, G. Peters, M. Hoffmann, P. Rudisch, L. Jakel, G. Lauckner, A. Grodrian, U. Guth, Continuous long-term monitoring of ruminal pH, Sens. Actuators B: Chem. 144 (2010) 395–399. https://doi.org/10.1016/j.snb.2009.01.012

[26] J. Shen, W. Zheng, Y. Xu, Z. Yu, The inhibition of high ammonia to in vitro rumen fermentation is pH dependent, Front. Vet. Sci. 10 (2023) 1163021. https://doi.org/10.3389/fvets.2023.1163021

[27] A.A. Megahed, P.D. Constable, Technical note: Evaluation of a colorimetric point-of-care test for measuring urine ammonium concentration in periparturient dairy cattle, J. Dairy Sci. 103 (2020) 8655–8660. https://doi/10.3168/jds.2020-18534

[28] M.D. Flythe, B.E. Harlow, G.E. Aiken, G.L. Gellin, I.A. Kagan, J. Pappas, Inhibition of growth and ammonia production of ruminal hyper ammonia-producing bacteria by Chinook or Galena hops after long-term storage, Fermentation 3 (2017) 68. https://doi.org/10.3390/fermentation3040068

[29] A. Adeniyi, I. Bello, T. Mukaila, E. Monono, A. Hammed, Developing rumen mimicry process for biological ammonia synthesis, Bioprocess. Biosyst. Eng. 46 (2023) 1011–1020. https://doi.org/10.1007/s00449-023-02880-7

[30] V. Lind, M.R. Weisbjerg, G.M. Jørgensen, J.E. Fermendez-Yepes, L. Arbesú, E. Molina-Alcaide, Ruminal fermentation, growth rate and methane production in sheep fed diets including white clover, soybean meal or Porphyra sp., Animals 10 (2020) 79. https://doi.org.10.3390/ani10010079

[31] H. Sun, K. Xiao, Z. Zheng, B. Yang, H. Duan, H. Zhao, Y. Zhang, Electroactive biofilm-based sensor for volatile fatty acids monitoring: A review, Chem. Eng. J. 449 (2022) 1–17. https://doi.org/10.1016/j.cej.2022.137833

[32] A. Korent, S. Trafela, K.Z. Soderznik, Z. Samardzija, S. Sturm, K.Z. Rozman, Au-decorated electrochemically synthesized polyaniline-based sensory platform for amperometric detection of aqueous ammonia in biological fluids, Electrochim. Acta 430 (2022) 141034. https://doi.org/10.1016/j.electacta.2022.141034

[33] K. Minami, H. Kobayashi, M. Matoba, Y. Kamiya, S. Maji, T. Nemoto, M. Tohno, R. Nakakubo, G. Yoshikawa, Measurement of volatile fatty acids in silage through odors with nanomechanical sensors, Biosensors (Basel) 13 (2023) 152. https://doi.org/10.3390/bios13020152

[34] J.J. Lamb, O. Bernard, S. Sarker, K.M. Lien, D.R. Hjelme, Perspectives of optical colourimetric sensors for anaerobic digestion, Renew. Sustain. Energy Rev. 111 (2019) 87–96. https://doi.org/10.1016/j.rser.2019.04.083

[35] T. Murata, K. Minami, T. Yamazaki, T. Sato, H. Koinuma, K. Ariga, N. Matsuki, Nanometer-Flat DNA-Featured Thin Films Prepared via Laser Molecular Beam Deposition under High-Vacuum for Selective Methanol Sensing, Bull. Chem. Soc. Jpn,  96 (2023) 29–34, https://doi.org/10.1246/bcsj.20220303

[36] R.G. Shrestha, J. Kawakita, Continuously activated function of aluminum in galvanic micro arrays in contact with water, J. Electrochem. Soc. 170 (2023) 080502. https://doi.org/10.1149/1945-7111/aceab3

[37] M. Mekawy, H. Noguchi, J. Kawakita, Quantitative and qualitative studies for real monitoring of interfacial molecular water, J. Colloid Interface Sci. 613 (2022) 311–319. https://doi.org/10.1016/j.jcis.2022.01.025

[38] E. Hajnal, L. Kovács, G. Vakulya, Dairy cattle rumen bolus developments with special regard to the applicable artificial intelligence (AI) methods, Sensors 22 (2022) 1–16. https://doi.org/10.3390/s22186812

[39] R.G. Shrestha, Y. Kubota, Y. Sakamoto, J. Kawakita, Quick and sensitive detection of water using galvanic-coupled arrays with a submicron gap for the advanced prediction of dew condensation, Sensors (Switzerland) 20 (2020) 1–10. https://doi.org/10.3390/s20113314

[40] L. Kumara, I. Rawalb, A. Kaurb, S. Annapoorni. Flexible room temperature ammonia sensor based on polyaniline, Sens. Actuators B: Chem. 240 (2017) 408–416. http://dx.doi.org/10.1016/j.snb.2016.08.173

[41] A. Kumar, B.K. Tudu, S.K. Pandit, Development of novel anti‐wetting coating on cellulosic surface using low carbon butyric acid, Cellulose 28 (2021) 4825–4834. https://doi.org/10.1007/s10570-021-03797-8

[42] M. Ruan, J.W. Wang, Q.L. Liu, F.M. Ma, Z.L. Yu, W. Feng, Y. Chen, Superhydrophobic and anti-icing properties of sol–gel prepared alumina coatings, Russ. J. Non-Ferr. Met. 57 (2016) 638–645. https://doi.org/10.3103/S1067821216060122

[43] C. Schanzenbach, D. Ilver, J. Blomgren, C. Jonasson, A. Krozer, L. Ye, B.O. Rustas, Preparation and characterisation of a sensing system for wireless pH measurements in vivo, in a rumen of a cow, Sens. Actuators B: Chem. 242 (2017) 637–644. https://doi.org/10.1016/j.snb.2016.11.089

[44] K. Yamamoto, N. Nishi, Hydrophobic hydration and hydrophobic interaction of carboxylic acids in aqueous solution: Mass spectrometric analysis of liquid fragments isolated as clusters, J. Am. Chem. Soc. 112 (1990) 549–558. https://doi.org/10.1021/ja00158a011

[45] M. Miura, Y. Nonomura, Evaluation of hydrophobicity for fatty acids using reversed-phase thin layer chromatography, J. Oleo. Sci. 68 (2019) 665–670. https://doi.org/10.5650/jos.ess19016

[46] R.P. Young, Infrared spectroscopic studies of adsorption and catalysis. Part 3. Carboxylic acids and their derivatives adsorbed on silica, Can. J. Chem. 47 (1969) 2237–2247. https://doi.org/10.1016/0021-9517(71)90096-0

[47] G.J. Gittens, Variation of surface tension of water with temperature, J. Colloid Interface Sci. 30 (1969) 406–412. https://doi.org/10.1016/0021-9797(69)90409-3

[48] J.P. Cloarec, C. Chevalier, J. Genest, J. Beauvais, H. Chamas, Y. Chevolot, T. Baron, A. Souifi, pH driven addressing of silicon nanowires onto Si3N4/SiO2 micro-patterned surfaces, Nanotechnology 27 (2016) 295602. https://doi.org/10.1088/0957-4484/27/29/295602

[49] K.W. Goyne, A.R. Zimmerman, B.L. Newalkar, S. Komarneni, S.L. Brantley, J. Chorover, Surface charge of variable porosity Al2O3(s) and SiO2(s) adsorbents, J. Porous Mater. 9 (2002), 243–256. https://doi.org/10.1023/A:1021631827398

[50] E. Virga, E. Spruijt, W.M. de Voas, P.M. Biesheuvel, Wettability of amphoteric surfaces: The effect of pH and ionic strength on surface ionization and wetting, Langmuir 34 (2018) 15174–15180. https://doi.org/10.1021/acs.langmuir.8b02875